## Slide 1: Gas Detection Training
MSR
Canary
Austech
Gastech

## Slide 2: Brief Intro to Units!
PPM
  Parts per Million, exactly what it sounds like!
  For example if you have 2g of CO2 in a 1kg bottle of air. How many parts per million of CO2 would be in the bottle?
PPK
  Parts per thousand is just parts per million * 1000.
  Similar example if you have 2g of CO2 in 1kg bottle of air. How many parts per thousand of CO2 would be in the bottle?
Additional note: PPM / PPK doesn’t actually measure parts it’s a volume fraction.

## Slide 3: Intro to Units %VOL
%VOL (Percentage Volume) (0-100% Vol)
Refers to the fraction of air that is occupied by the gas.
For example:
If 5% Vol CO2 is in the air, CO2 makes up 5% of the volume of air.
Note:
If 5% Vol of CO2 is in the air, it is equivalent to 50,000 ppm of CO2. This is a conversion valid for all gases (even though it is an estimation).
Exercise:
If a CO2 gas detector is reading 0.4%Vol in a room with a volume of 55m^3 what is the CO2 reading in ppm.

## Slide 4: Intro to Units % LEL
LEL is referred to as the lower explosive limit (AS 60079.2)
  Sometimes referred to as Lower Flammable Limit
UEL is referred to as the upper explosive limit (AS 60079.2)
  Sometimes referred to as Upper Flammable Limit
Concentration of Explosive Gas Vs Air

## Slide 5: Lower Explosive Limits (AS 60079)
Lower explosive limits are defined by AS 60079.20.1-2012
Methane:
4.4% VOL = 100% LEL
Hydrogen:
4.0% VOL = 100% LEL

## Slide 6: IECEx vs. ATEX
IECEx is a certification that requires adherence to IEC/AS 60079.
IECEx requires recertification and quality management in manufacturing to the point where recertification every few years is required. ISO Type 5 certification this means spot checks for conformance is performed and any changes to manufacturing and end product require recertification.
ATEX is not a ISO Type 5 certification, thereby once certification is acquired the manufacturing conformance and end product modifications are not verified by certification body.

## Slide 7: Hazardous Area Standard AS 60079 (Note: Gas Only)
Hazardous gas sensors are made for 3 types of zones (IEC 60079-10.1). Each reflecting the maximum likelihood of flammable gas.
Zone 0: Explosive gas atmosphere present 1000 hrs / year. (>10% of time in service)
Zone 1: Explosive gas atmosphere present 10-1000 hrs /year (1-10% of time in service)
Zone 2: Explosive gas atmosphere present <10 hrs / year (0-0.1% of time in service)
Note: Boundaries and zones should come clearly labelled by consultants.
Most often they are not clearly labelled.

## Slide 8: Hazardous Area Standard AS 60079 Cont. (Selection of Equipment)
Few things to consider:
Minimum Ignition Energy
Auto Ignition Temperature (VERY IMPORTANT)
Type of gas being detected.
T Rating (Found by using the Auto Ignition Temperature)
Equipment Protection Levels (Zone 0 / Zone 1 / Zone 2)
IP rating required (if explosion protected IP rating shall be sufficiently high)

## Slide 9: Hazardous Area Standard AS 60079 Cont. (Selection of Equipment)
Auto Ignition Energy is a measure of the maximum temperature allowed in the hazardous area.
For Example CS2 (carbon disulfide) has a auto ignition temperature of 85 C, which means if the surface temperature of any equipment in the room reaches 85 C there is a danger of spontaneous combustion.
Using the auto ignition temperature in AS 60079.20 & 60079.14 we can select the appropriate gas sensor & conduit / wire to use in the zone rated area.
Note: When selecting the sensor (given no T class rating) take into consideration the ambient temperature (Default ambient is -20 to 40C ). If ambient temperature in install location is higher adjust T class to compensate. Furthermore for a specified auto ignition temperature between 300-450 C therefore a Temperature class of T2-T6 is appropriate.

## Slide 10: Hazardous Area Standard AS 60079 Cont. (Selection of Equipment)
Gas Group, Gases are classed into 3 categories A, B , C.
These can be determined using AS 60079-20 by looking up the gas in Annex B.
Ignition energy is the MIE shown.
If sensor or equipment requirement is IIA which of the sub-divisions are suitable?

## Slide 11: Hazardous Area Standard AS 60079 Cont. (Equipment protection level)
Types of Protection
Flame proof Enclosure {Ex d}
Flameproof enclosures ensure that any flame generated by electronics inside of the housing cannot spread to the outer atmosphere.
Increased Safety Enclosure {Ex e}
Increased safety specification inhibits arcs or sparks from forming in normal operation. Generally used for junction boxes or lighting fittings.
Intrinsic Safety {Ex i}
A circuit designed in order to produce sparks or thermal effects which lie below the minimum ignition energy.

## Slide 12: Hazardous Area Standard AS 60079 Cont. (Equipment protection level)
Similar designation for Ex da/db/dc and so on.
Note:
G stands for Gas
D stands for Dust
EPL:
Equipment Protection Levels are designated by the hazardous goods consultant. These should be supplied as either a zone or a EPL value as shown in table below.
Equipment Classification:
Used as a method of differentiating the EPL along with protection method.

## Slide 13: Hazardous Area Standard AS 60079 Cont. (Putting it all together)
Exercise:
Determine a suitable component to meet the specifications below:
Zone 1
Auto Ignition temperature of 365C
Hydrogen gas detector (Gas Group B)
Ambient temperature = -30 to 60C
Gas Detector A
EX db IIB T5
Gas Detector B
EX da IIC T6
Gas Detector C
EX da IIC T1

## Slide 14: Types of Sensors – Electrochemical sensors
Electrochemical sensors are similar to batteries and have two or sometimes 3 electrodes.
They work on a similar principle where if gas is present the current or voltage between the pins changes. This can be measured in (microVolts or NanoAmps)
This isn’t really useful but, its nice to flex every now and then.
Things to remember:
These sensors are not very gas specific (Check for cross sensitivities and account for them)
Require a warmup time of 1-2 hrs (Oxy generally 2+ hrs)
Primary failure occurs due to power fluctuations & the drying of electrolyte which occurs in hot weather.
Lifespans (O2 – 3yrs, CO – 5 yrs … etc consult datasheet)
Calibration: 6 months (Oxy) to 1 yr (Carbon Monoxide)

## Slide 15: Types of Sensors – Pellistor or Catalytic sensors
Catalytic bead or Pellistor sensors work on the principle that the gas being detected is flammable.
Things to remember:
Not very gas specific (methane sensors will detect hydrogen and vice versa)
Works on almost all explosive gases, however some special explosive gases require special filters, which means the regular sensorcells will not work on these types of gases and a specific sensor is required.
Pellistor sensors can be directly affected by the humidity and water splashes. This may result in a sensor reading higher than ambient gas levels.
Primary failure occurs due to silicon poisoning. DO NOT USE SILICON ANYWHERE NEAR THESE!!!!
Lifespan: 5+ years
Calibration Requirements: 6 monthly calibration

## Slide 16: Types of Sensor – Infra-red
Infra-red Sensor detect gases because CO2 and other gases can absorb Infra-red light at a certain wavelength. Due to this specific property only certain gases are available with infra-red detection cells.
Things to remember:
Infra-Red sensors are very specific and very accurate.
Primary failure will occur due to dust ingress or electrical circuit failure due to power spikes.
Lifespan: 7+yrs (This value is conservative, it’s likely longer)
Calibration: Explosive infra-red (1 YR), CO2 (1 YR)
Bump test: Every 6 months (check whether they respond)

## Slide 17: Calibration Tool Kit Essentials
Tools Required:
Essentials:
Gas Regulator
Sensor cap (to ensure good seal between calibration gas and sensorhead)
Flexible hose (2m length 8mm OD)
Multimeter with leads for alligator clips
Gas Bottle 1:
18.5% Vol Oxygen
Bal. N2
Gas Bottle 2:
Mix Gas:
2.5% Vol CH4
100 ppm CO
18% Vol Oxygen
4% Vol CO2
Bal. Air / N2
Gas Bottle 3:
99.999% Vol N2

## Slide 18: Calibration of MSR Infra-red & Electrochemical Sensors
Step 1: Engage Service mode (To inhibit alarm when calibration gas is applied)
Step 2: Apply multi-meter to Test Pins
Step 3: Apply Zero calibration gas to sensor (wait 30 seconds)
Step 4: Check multi-meter reading. At Zero gas value it should read 40mv. If it requires adjusting Pull board out and plug back in. Hit the zero button multiple times until the multi-meter is reading 40mv.
Step 5: Remove zero calibration gas and apply Gain calibration gas (wait 30 seconds, Consult datasheet for t90 to decide on appropriate waiting time)
Step 6: Adjust gain calibration pot till given value is equivalent in mv to the following equation.

## Slide 19: Calibration of MSR Pellistor sensors
Step 1: Engage Service mode (To inhibit alarm when calibration gas is applied)
Step 2: Apply multi-meter to Bridge Pins
Step 3: Apply Zero calibration gas to sensor (wait 1 minute for sensor reading to stabilize).
Step 4: Adjust Bridge Voltage to 0 mv (slightly above 0 mv is recommended ~0.1mv).
Step 5: Apply multi-meter to test pins.
Step 6: At Zero gas value it should read 40mv. If it requires adjusting Pull board out and plug back in. Hit the zero button multiple times until the multi-meter is reading 40mv.
Step 7: Remove zero calibration gas and apply Gain calibration gas
Step 8: Adjust gain calibration pot till given value is equivalent in mv to the following equation.

## Slide 20: Adjustment of AUSTECH Sensor System
Accept
Reset
Select
Step 1: Cycle to sensor to be calibrated
Step 2: Apply zero gas
Step 3: Enter Calibration mode by using the Accept and Select Buttons together for 5 seconds. (Countdown shown shall be CAL 5, 4, 3… 1 if LCD displays SET, abort and ensure accept and select buttons are used.)
Step 4: Adjust sensor reading on LCD to 0 reading.
Step 5: Apply gain gas
Step 6: Adjust sensor reading on LCD screen to match the gain calibration gas concentration.

## Slide 21: Calibration of AUSTECH Sensor System
Step 1: Apply zero gas and adjust sensor output using zero pot
Step 2: Apply gain gas and adjust sensor output using gain pot
Sensor
Sensor

## Slide 22: Calibration of AUSTECH Sensor System Continued…
mA Current Generator
Step 1: Apply 4ma through the current generator
Step 2: Scroll to sensor being calibrated and hold down “Accept” and “Select” buttons
Step 3: Adjust sensor reading to 0
Step 4: Apply 20 ma through the current generator
Step 5: Adjust sensor reading to full range (ie. 25% for oxygen or 100% LEL for methane)
Note: Find the connection points on HOSLAB Server – “L:\Manuals\Austech\”

## Slide 23: OLDHAM OLCT-10N Auto-Calibration
Cal ON refers to applying the calibration cap below

## Slide 24: OLDHAM OLCT-10N Auto Calibration – Cont.
Calibration Gas Concentration can’t be adjusted. If you don’t have the correct gain gas perform manual calibration (bit more effort)
Common Sensor Replacements (Suitable types):
Alphasense O2-A2 only work for this
If LEL sensor, replace the entire sensor as no replacement sensor cell available

## Slide 25: OLDHAM OLCT-10N Sensor Replacement
Step 1
Step 2
Cable in cal kit – 2x bought by Hoslab.
Step 3
Apply zero gas. Adjust Pot labelled ZERO until multi-meter reads 0 mv. DON’T USE THE MAGNETIC CAP FOR THIS!!
Step 4
Apply Gain gas. Adjust Pot labelled “SENS” until multi-meter reads the correct signal value in (mv) according to formula below. DON’T USE THE MAGNETIC CAP FOR THIS!!
Step 5
Perform Auto Calibration with magnetic Calibration cap!

## Slide 26: Dynament MSR Sensor
Negative / Positive zero suppression:
These sensors drift and therefore they require a suppression function which enables the sensor to output 0ppm until it crosses the set suppression threshold. For example if threshold set at 10 ppm, sensor will read 0 ppm from 0-10 ppm.
Negative suppression works similarly in that until it goes below say -10 ppm in output the sensor will show 0 ppm.
Error Codes flash on normal LCD screen replacing reading for 2 seconds.
Menu
Enter
Next
Prev
Code # | Issue
2 | Negative drift, adjust negative suppression
4 | Zero the sensor as the sensor is too far outside of negative suppression
8 | Sensor needs to be gain calibrated
16 | Calibrate sensor, if error still occurs replace sensor cell if possible.
32 | Calibrate sensor, if error still occurs replace sensor cell if possible.
64 | Sensor signal too high, attempt calibration and replace if it doesn’t work
99 | Config. data corrupt, attempt power cycle and recalibration if possible

## Slide 27: Dynament MSR Negative Suppression
Step 1:
Press Menu key navigate to option E:20 using next /prev keys and press enter.
Step 2:
Use the next and prev keys to increase and decrease the negative suppression (0-10% OF RANGE)
Step 3:
Press enter to save the negative suppression value.
Menu
Enter
Next
Prev

## Slide 28: Dynament MSR Positive Suppression
Step 1:
Press Menu key navigate to option E:19 using next /prev keys and press enter.
Step 2:
Use the next and prev keys to increase and decrease the negative suppression (0-10 % OF RANGE)
Step 3:
Press enter to save the negative suppression value.
Menu
Enter
Next
Prev

## Slide 29: Dynament MSR Calibration
Step 1:
Press Menu key navigate to option E:1 using next/prev and press enter
Step 2:
Apply zero gas for 30sec to 1 min. Press enter to zero.
“- - - -” WILL BE DISPLAYED IF ZERO IS SUCCESSFUL
Menu
Enter
Next
Prev

## Slide 30: Dynament MSR Calibration cont.
Step 1:
Press Menu key navigate to option E:2 using next/prev and press enter
Step 2:
Apply gain gas for 30sec to 1 min.
Step 3:
Use next/ prev buttons to increase / decrease the reading till the concentration reaches the gain gas concentration.
Step 4:
Press enter to complete calibration.
“- - - -” WILL BE DISPLAYED IF GAIN CALIBRATION IS SUCCESSFUL
Menu
Enter
Next
Prev

## Slide 31: Sensor Info

## Slide 32: MSR MC2 “Analog” Sensor Calibration
Step 1: Apply Service mode (Inhibit alarms on the controller)
Step 2: Wait for power on time / warm-up Time to finish and navigate to
“Installation & Calibration” -> “Calibrate” -> “DP 1”
Step 3: Apply zero gas to sensor (wait 30 seconds) & go to “Zero calibration” menu.
Step 4: Hit enter , if calibration is accepted “Saved” will appear on screen.
Step 5: Navigate to “Test gas” and enter “Gain DP 1”
Step 6: Hit enter , if calibration is accepted “Saved” will appear on screen.

## Slide 33: MSR SC2 “Digital” Tool Operation
Plug the PC tool into your laptop.
USB drive should be detected, if not call ALVI / Sandeep for troubleshooting.
Open PCE06-XXX-XXXX.exe file
Enter Password for software
“8304”
Select Comm port
Click Connect

## Slide 34: MSR SC2 “Digital” Sensor Calibration
Step 1:
Before making any changes special mode or service mode must be activated
Step 2:
Engage Calibration Mode by clicking this icon.
Step 3:
Fill In Calibration Details and select sensor.
Step 4:
Hit the arrow key at top right corner
Start Zero Calibration:
Apply zero cal gas
Click Display
Wait 30 seconds
Click Calculate
Click save once sensor reading has stabilised

## Slide 35: MSR SC2 “Digital” Sensor Calibration cont.
Step 5:
Hit the arrow key at top right corner to progress on to gain calibration.
Enter Bottle Gas Concentration.
Apply gas
Click Display -> Calculate -> Save
Allow enough time for calculation, some sensors require upwards of 1 minute in order to complete

## Slide 36: Sensitron Sensor Calibration (Tool Required)
Step 1:
After you plug in the calibration keypad, after displaying the software version. The module will display the enter password menu.
“0000”
Using up and down enter pass: “0261 “
Step 2:
If password entered correctly “Cal” will be displayed.
Hit Enter
“Zero” will be displayed

## Slide 37: What is the difference between Digital and Analog?
Lets list some advantages and disadvantages
  Digital Sensors!
  Very Accurate
  Very easy to calibrate
  Figuratively cannot fuck it up! 
  Most modes of failure can have an error code ie if there’s an error in the sensor cell it will give a different error code to comms failure.
  Analog Sensors!
  Difficult to calibrate, require adjustment of input and output
  Multiple points of failure only one level of response in terms of error reporting sensor just goes underrange when in error.

## Slide 38: PG2 Topics Covered
SC2 & MC2 calibration
  SC2 & MC2 Calibration
  SC2 & MC2 Troubleshooting
Common Nomenclature
Easy Config Basic UI explanation
  Logging
  Controller config files
  Types of Bus Connections explanation(L BUS , F BUS, M BUS)
  Calibration
GC-06 Config
  Parameters Config
  Analog Output / Analog input calibration
  Troubleshooting Errors
  Wiring Configurations for BUS
SB2 & MSB2
  Addressing SB2 + PARAMETERS FOR DGC BUS
  Addressing MSB2 + PARAMETERS FOR DGC BUS
  Addressing SB2 for Modbus output
  Addressing MSB2 for Modbus output
  MSB2 Analog output (Remote)
EP-06 Config
  Parameters config
  Analog Output (Remote)
  Troubleshooting Errors
Modbus Setup
  Modbus Version Check: How to
  Modbus Addressing
  Technical Specifications of Modbus output
    Baud, RS 485
Modbus Poll Verification, How to:
BacNET GATEWAY addressing
DEM-06 Config and Features

## Slide 39: Common MSR Sensor System Error Codes
Error Code | Hints
9000 | Wiring issue or communication issue between controller and sensor or sensor and sensor board.
8001 | UPS Issue, check the voltage being delivered from the power supply, ensure it is below 25 VDC
8100 | Sensor is under range Calibration issue (just zero and gain calibrate again) Check programming to make sure you’ve programmed the correct sensor (digital / analog)
8200 | Sensor is over range Calibration issue (Sometimes LEL sensors go over range, just perform a zero calibration and gain calibration) Power issues can cause over range signal due to the fuse resistor being burned out.
8002 | Sensor ADC error Sometimes the controller will show 0X8200 error and if you plug into the sensor it will show 0x8002. This is a warranty issue, replace and move on.
8001 | Sensor EE error Sometimes caused by programming (double check) otherwise replace.
8020 | I/O Error -Programming / Analog output interface error.

## Slide 40: SC2 & MC2 Sensor Calibration
Yellow light flashing
  Fault / Power on time
    Best way to check is to read measuring values
Green light
  Power on
Red Light
  Alarm situation
MC2 Handtool LED Indicators
Basic Overview
Insert Tool into SB2 OR MC2 PINS
Wait for Power on time to finish
Zero calibration first
Gain Calibration afterwards
Sample Power on Times:
CO2
  30 Seconds
Refrigerant
  180 Seconds
CO
  8 Seconds
Oxygen
  10 Seconds

## Slide 41: SC2 & MC2 Troubleshooting
Reading Operating Data
  Must Use EASY Config tool for this
  Connect SC2 OR MC2 to SB2 OR MSB2
  Read Operating Data (This gives you information on usage of sensor)
    Max Operating Temp
    Min Operating Temp
    Can Use this info to determine warranty outcome
Common Wiring Issues
  Overrange
    Incorrect Wiring
  Underrange
    Requires re-calibration (Most prevalent in LEL)
Common Troubleshooting methods:
Incorrect wiring for long periods can cause damage to Analog board on MC2.
3 pin MC2 connection does not work for MC2 HAND TOOL then sensor is dead.

## Slide 42: Common Nomenclature
AP = ANALOG SENSOR MEASURING POINT
DP = DIGITAL SENSOR MEASURING POINT
AR / PR = Relays inside control cabinet
SR = Signal Relays in the field on field devices
SB = Sensor board
MSB2 = Multi Sensor Board
MSC2 = Multi Sensor Controller
MGC2 = Multi Gas Controller
SC2 = Digital Sensor head
MC2 = Analog Sensor head
PG2 = Polygard 2
PX2 = PolyXeta 2 Hazardous Area system
AO = Analog Output
FBUS = Field Bus
MBUS = Main Bus
LBUS = 3 wire bus
MODBUS = Industry standard comms developed by Schneider over IP/RS485
BACNET = Industry standard comms over IP/RS485
BAUD RATE = baud rate is the rate at which information is transferred in a comms standard
Comm port = Communication port on your computer

## Slide 43: Types of Comms Native to PG2 SYSTEMS
Field bus
  What communicates on this bus?
    EP
    SB
    MSB2
    MSC2
    MGC2
  Special Features
    DGC-06 proprietary bus
    4 wire bus or 2 wire RS485 assuming common ground
    Baud rate 19200
    Based on MODBUS
L-BUS
  What communicates on this bus?
    SC2
    Remote board
  Special Features
    3 Wire bus Not much known by us about this bus
    Suspect it is MODBUS 3 wire
M-BUS
  What communicates on this bus?
    GC-06 ONLY!
  NOT READY YET!!
  Special features (yet unknown)
    Communications between two separate systems which can then communicate via HLI to Modbus only requiring connection to 1 system ( refer to Fig 1.1).
    Most likely a 4 wire bus
SYS1
SYS2
SYS4
SYS3
HLI connection to Building Management System
M-BUS connection between systems
Fig 1.1 (M-Bus example)

## Slide 44: DGC-06 HLI Outputs available
BACNET
  BACNET Modules communicate with GC-06 on a MODBUS connection therefore two things are required for BACNET Modules to work. A Modbus capable controller + BACNET Module
  Limitation:
    Because of address limitations BACNET module can only output the following:
Modbus
  Easiest and most cost effective solution for HLI comms output for our system to BMS.
  Limitations:
    None
      All system parameters are available for viewing via Modbus output.
Fig 1.2 (BACNET Module configs)

## Slide 45: Easy Config Basic UI
Tool Number: Please note this should match the physical tool you’ve been provided.
Connect / Disconnect: Connects / disconnects to device that is plugged into tool.
Digital Sensors + Signal Relays 1-96
Analog Sensors + Analog Relays 1-32
Analog Outputs from GC/SB/MSB/EP
Datalogging Toolbar
Full explanation later
Operating Data
System Parameters
Relay Parameters
Measuring Parameters
Addressing Menu
Fig 1.3 (Easy Config UI)
Software Ver. Of Device Connected
Comm Port used to connect & Baud Rate
Connection status Green = good

## Slide 46: Measuring Parameters
Measuring Parameters
Very Similar to OLD DGC-05 easy config page.
New Features Highlighted in fig 1.4
Fig 1.4 MP CONFIG
Fault function for Alarm
Eg: If fault exists Alarm 1 ON MP XX will trigger relay XX
Latching Function works differently in GC-05.
Eg: If alarm 4 latching, then automatically Alarm 1-4 are all latching
Toggle to select AR / SR (digital or analog relay)
Toggle to select AR / SR (digital or analog relay)

## Slide 47: Relay Parameters
Relay Parameters
Very Similar to OLD DGC-05 easy config page.
New Features Highlighted in fig 1.5
Fig 1.5 Relay CONFIG
Activate by Failure (Unknown or unimplemented)
Activate by Maintenance (Unknown or unimplemented)
Alarm Voting Quantity:
EG: Relay will only activate if 2 or more (based on number quantity input) separate alarms assigned to this relay are in alarm.

## Slide 48: System Parameters
System Parameters
Very Similar to OLD DGC-05 easy config page.
New Features Highlighted in fig 1.6
Fig 1.6 System CONFIG
Activate Failure from (Unknown or unimplemented)
Maintenance interval function not implemented.
Special Function:
Title 24 / Default
Title 24 is a US regulation, that is not useful to us.
Error Time (Sec) not implemented yet.

## Slide 49: Logging setup
Logging setup
Very Similar to OLD DGC-05 easy config page.
New Features Highlighted in fig 1.7
Fig 1.6 System CONFIG
Logging options:
Digital Sensors
Analog Sensors

## Slide 50: Wiring Configurations for F-BUS

## Slide 51: Practical Exercises
Programming Sensors
Programming Relays
Calibrating Sensors
Analog Output programming
Analog Output Programming GC TO EP
Analog Output Programming SB TO GC
Analog Output Programming SB TO MSB2
Logging Sensors
Hand tool calibration
  PG2 Handtool
  MC2 Handtool
PX2 Sensor calibration
PX2 Sensor Relay programming
Modbus Addressing
Modbus version check
Modbus Poll Check